## Slide 1: An Inductive-Capacitive (LC) Sensor for Real-time Biofilm Growth Monitoring
Ekaterina Tolstaya
Young Kim

## Slide 2: Biofilm Sensing
Medical biofilms
  Cause diseases, often requiring surgery or antibiotics at does 1000 times great than for non-biofilm infections
  Biofilms forming in catheters can cause infections in patients
Environmental biofilms require cleaning with harsh chemicals
Biofilm formation on the solid surfaces of the catheter within the body of the patient.

## Slide 3: Why Inductance - Capacitance Sensing?
Ultimate Goal – Flexible biofilm sensor inside of a catheter
Young’s SAW sensor cannot be used on flexible substrate due to changes in piezoelectric coefficient
Further work – Wireless integration using inductive coupling
Ease of fabrication
  2 steps vs. 7 steps for SAW sensor
  No piezoelectric material deposition (ZnO)
Photolithography mask for the LC sensor

## Slide 4: Principles and Design of the LC Sensor
Considerations
  Increase sensitivity
  Keep resonance frequency within measurable range
  Ease fabrication
Biofilm Growth
Permittivity Constant Change
Capacitance Change
Resonance Freq. Change
Parameters
  Device radius
  Electrode thickness
  Electrode gap
  Electrode width
The capacitance of the sensor depends on the permittivity constants of the medium and SiO2.

## Slide 5: My Role
Simulations
Mask design
  Automated of drawing spirals in AutoCAD
Fabrication
  Learned photolithography, lift-off, inspection of device
Test setup
  Designed a modular system to ease cleaning and containment of bacteria
Experiments
  Confirmed device operation
  Conducted biofilm experiments
Developed negative resist
Lift-off process
Assembled test setup and sensor

## Slide 6: Device Parameters
Device | Radius | Electrode Width | Electrode Gap | Inductance | Capacitance | Simulated Resonance Freq. in Water | Actual Resonance Freq. in Water
Small | 0.65 cm | 60 um | 80 um | 0.43nH | 1.2nF | 221.56MHz | 136 MHz
Large | 1.5 cm | 60 um | 80 um | 8.57nH | 8.77nF | 18.36MHz | 14.7 MHz
Photolithography masks for the large and small LC sensors
400 um
800 um

## Slide 7: Circuit Simulation
Matlab parameter simulation:
  Capacitance = 1.2 nF
  Inductance = 0.43 nH
  Large load resistance due to instrument impedance
PSPICE circuit simulation:
  Circuit will act as a band-pass filter, with a resonance frequency of 221.3 MHz
LC Sensor simulated in PSPICE as an parallel LC circuit

## Slide 8: Device Fabrication
Negative Resist
[Image reversible]
AZ5214
Cr/Au
(15 nm/200nm)
By E-Beam Evaporation
Silicon Oxide using PECVD
Photolithography mask for the LC sensor
200 um
100 um
Microscope photo of the coils of a fabricated sensor
Lift-off
Notes: SCALE BARS

## Slide 9: Test Procedure and Equipment
Test setup - modular, reusable and resistive to bleach clean, with a 80 mL volume
Escherichia coli W3110 in LB media
Network analyzer as frequency sweeper and spectrum analyzer to analyze frequency response
Device interface is controlled through GPIB by C code
1)
2)
3)
4)
Time
Frequency

## Slide 10: Characterizing Device Sensitivity
Average resonance frequency:
  Air: 15.6 MHz
  Water: 14.7 MHz
Statistically significant difference between air and water measurements
High noise because of low capacitance
Sensitivity = 1140 Hz/delta dielectric
Chart: Confirmation of Device Operation
| Air | Water
 | 15649000 | 14747000

## Slide 11: Sensing Biofilm Growth
Frequency observed to increase by 0.25 MHz as biofilm grows
Consistent with literature
  Water permittivity: ~80
  Biofilm permittivity: ~60
Filtered to reduce noise
  Moving average of 400 samples

## Slide 12: Additional Biofilm Experiment
Different device
Shorter time frame
Same filtering method

## Slide 13: Future Experiment Ideas
Determining actual change in permittivity during biofilm growth
Imaging biofilm on sensor after biofilm experiments
Summarizing simulation data
  Simulating change in capacitance over time – PSPICE, Matlab
Characterizing permittivity change by drying biofilm on sensor
  Compare permittivity for dry clean sensor and sensor covered in dry biofilm